Section E, Rev. 3, 11/25/03
Tek-CARE® NC150N—Audio-Visual Nurse Call System


Engineers’ Specification

A.
The contractor shall provide and install all equipment, accessories and materials necessary for a complete operating system in accordance with the following specifications.

B.
The equipment furnished under this specification shall be the standard product of one manufacturer and shall be as manufactured by TekTone® Sound & Signal Mfg., Inc.

C.
The equipment to be installed must comply with all ap​plicable codes and regulations including (but not limited to) NFPA 70, 99. Each piece of equipment to be installed as part of the nurse call system must be listed under the current version of Underwriters Laboratories Inc. standard UL®1069.
D.
The contractor shall guarantee availability of local service and available stock of the manufactures standard parts.

E.
The contractor shall be responsible for providing a complete functional system including all necessary components, whether included in this specification or not.

Part 2—System Configuration and Operation

2.1 Configuration

The system shall consist of bed stations, emergency stations, dome lamps, annunciators, duty stations and power and control equipment, all as specified herein and shown on the plans.

2.2 Operation

A.
Routine Call—A routine call shall be originated at the bed station by pressing the momentary button on the call cord. This shall cause the following:

1.
The call placed light on the bed station shall illuminate.

2.
The dome light associated with the calling station shall light steady white.

3.
The zone light associated with the calling station shall light steady white 

4.
The annunciator lamp/button and normal call lamp shall light steady on the nurse master panel.

5.
A slowly pulsing audible tone shall sound at the Nurse’s station and Duty stations if provided.


Pressing the tone-off button shall silence the routine call tone. A routine call may be canceled at the nurse’s master station or at the calling station. To answer the call, the appropriate selector lamp/button is depressed and the operator may talk via the speaker and talk button or by the optional handset. Any system requiring the use of a button with the handset shall not be acceptable. In the event that a call cord is removed from the patient station, a call shall be placed at the annunciator.

B.
Emergency Call—An Emergency call shall be initiated at any emergency pull-cord station by pulling the cord or pressing down the large lever on the station. This shall cause the following:

1.
The call placed light on the emergency station shall illuminate and flash.

2.
The dome light associated with the room shall illuminate and flash.

3.
The zone light associated with the room shall illuminate and flash.

4.
The annunciator lamp/button and emergency call lamp shall illuminate flash on the annunciator panel.

5.
A tone signal shall sound in one-second intervals.


An emergency call cannot be canceled at the master. It must be canceled at the calling station. Systems allowing the reset of an emergency call other than at the point of origin shall not be acceptable.

C.
Staff Follower—It shall be possible to forward calls to a pre-designated room. By pressing the appropriate station selector button and then the "Duty" button, all calls on the system will automatically be forwarded. The tone rate at the receiving station shall equal the priority of the call.

Part 3—Equipment

3.1 Bed Stations

A.
The bed station shall contain a call-placed indicator, a 1/4" jack for a call cord, a 3.5" 45 ohm speaker, a cancel button and a non-private indicator mounted on an attractive ABS plastic faceplate. Plug-in connectors shall be provided for wire connections. The bed station shall be TekTone® IR151B.

B.
The dual bed station shall operate like the IR151B, but will include two call-placed indicators and two 1/4" jacks. The dual bed station shall be TekTone® IR152B. 

C.
The pull-cord station shall operate like the IR151B, but will include a pull cord instead of a 1/4" jack for call placement. 

3.2 Emergency Pull Cord Switch

The emergency pull cord station shall contain a red call-placed indicator and 6’ nylon cord with pendant mounted on a single gang ABS plastic faceplate stamped EMERGENCY. The cord shall pull through an eyelet guide to ensure a 360-degree even pull. Pigtail connectors shall be provided for wire connections. The emergency pull cord station shall be TekTone® SF155B. The waterproof emergency switch shall be equal to the SF155B but shall use a magnetic switch and “O” rings. The waterproof station shall be TekTone® SF337C.

3.3 Staff Emergency Switch

The staff emergency station shall be the same as the emergency pull cord station. The pull cord shall not be installed and the 1” lever will be pulled down for call activation. The emergency pull down switch shall be TekTone® SF155B.

3.4 Code Switch

The Code switch shall contain a “call-placed" indicator mounted on a single gang, ABS white plastic plate with a large blue actuator labeled CODE. Color-coded pigtail wires shall be provided for wire connections. The code switch shall be TekTone® SF156B

3.5 Dome Light/Zone Light

The dome light shall contain a bulb with a translucent plastic cover mounted on a two-gang, plastic faceplate. The white cover shall be removable without the use of tools and the bulb shall be bayonet mount to allow ease of bulb replacement. Any dome light requiring the use of tools or disassembly for bulb replacement shall not be acceptable. The corridor dome light shall be TekTone® LI381. The zone light shall be the same as the LI381 but with an additional bulb with a red colored bulb cover. The zone light shall be TekTone® LI382.

3.6 Duty Station

The duty station shall contain a tone signaling device and a call indicator LED mounted on a single gang, ABS plastic faceplate. Pigtail connectors shall be provided for wiring connections. For a normal call, the LED shall light steadily and a tone will sound every 8 seconds. For an emergency call, the LED shall flash and a tone will sound every second. The duty station shall be TekTone® LI150B.

3.7 Annunciator

The annunciator shall contain individual lamp/button corresponding to each calling location. It shall be possible to configure the lamp/button in multiples of four to allow exact configuration. Additionally, it shall contain a tone signaling device, lamp test, call, duty (nurse follower) and talk push buttons, a 3.5" 45 ohm speaker, a voice switched handset and an integral amplifier. The indicator lamps shall have removable lenses for lamp replacement and to protect lamp identification markings. The annunciator shall be of modular construction with a metal, flush wall housing, extruded aluminum panel(s) and aluminum trim frame. Screw terminals shall be provided for wire connections. The annunciator shall be TekTone® NC150N.

3.8 Power Supply/Control Unit

The power and control unit shall be housed in an appropriate housing and shall provide for the call signal, tone silencing, and power circuits. Screw terminals shall be provided for wiring connections. The annunciator power supply/control unit and cabinet shall be TekTone® IH151NK. 

3.9 Third Priority Control Unit (Optional)

The third priority unit shall be of rugged aluminum construction mountable on a flat surface or appropriate electrical box (see IH151NK). Screw terminals shall be provided for wiring connections. The unit shall be used in addition to the PK151A and will supply a third priority tone and light signal for ancillary annunciation of smoke detectors. The third priority control unit shall be TekTone® PK153

3.10 Call Cords

The call cords shall be seven (7) or ten (10) feet, single or dual (“Y”) with a 1/4" plug, bed clip and momentary push button. Call cords shall be TekTone®, SF301 for single, SF302 for dual. Add the suffix “/10” for ten-foot models

3.11 Door Remote Stations:

It shall be possible to integrate standard door intercom stations to the NC150N system. The button module shall be PM151/4. This integration will allow the use of these door stations: OR501B flush mount metal, OR603 surface mount ABS plastic, OR501S surface mount metal.

Part 4—Radio Pager Interface Unit

The Voltage Interface Unit shall have the capability to accept analog inputs and convert them into serial data allow the interface to a radio-paging transmitter. The voltage interface will accommodate 128 inputs and is connected directly to the nurse call system. The interface unit shall be TekTone® NC375

Part 5—Radio Pager Transmitter Unit

The UHF on-site transmitter shall be specifically designed to receive patient calls from the NC150N/NC375. The unit shall transmit patient call information including at least the room number and priority of call in Alphanumeric format to TekTone®’s NC395 pocket pager. The unit shall transmit higher priority calls first regardless of chronological order received. The unit shall have a transmitting range of up to one mile (dependant upon results of a site survey) and shall support up to 9999 pagers. The radio pager transmitter shall be TekTone® NC366A.

Part 6—Wiring

Wiring shall be in strict accordance with TekTone® IL381 Installation Manual wiring diagrams.

Part 7—Warranty

Each piece of equipment shall be warranted by the equipment manufacturer to be free of defects in material and workmanship for a period of twelve (12) months from the date of shipment from the factory.

Part 8—Shop Drawings

The electrical contractor shall submit shop drawings to the electrical engineer for approval. The contractor must have approved shop drawings in his possession prior to installation.

Website Disclaimer and Legal Notice for downloadable 
Architects & Engineers Specifications

The specifications attached herein are being supplied in a text format for your convenience. TekTone® Sound & Signal Mfg., Inc. is not responsible for any changes made to the original documents.

As with any other TekTone® Sound & Signal Mfg., Inc. document, please be advised that the information was prepared by TekTone® Sound & Signal Mfg., Inc. as general information and a service to our customers. While we endeavor to supply accurate information, we can take no responsibility for errors or misinterpretations in any of the content of this site. We also reserve the right to remove or change literature and pages from this site where legal requirements are contravened or where complaints are received and upheld.

Except in the case of third-party material reproduced on this site, the information and all content, including logos and artwork, provided herein is protected under the U.S. and Foreign Copyright Laws and may not be copied, reproduced or retransmitted by any means, electronically or otherwise, without the express written permission of TekTone® Sound & Signal Mfg., Inc.

Modifications. All modified specifications will include the following ledger: Modified specifications are based upon TekTone® Sound & Signal Mfg., Inc. product specifications. Modified specifications have not been reviewed or approved by TekTone® Sound & Signal Mfg., Inc., unless so noted.

Indemnification. It is agreed that TekTone® will be held harmless from any claims, causes of action, pending or threatened, losses, awards, damages or attorney’s fees arising out of any suit or claim relating to the use and/or modification of the specifications.

Reservation of Rights. This disclosure is not intended, nor shall it be construed, to release any right, title or interest in any name, trademark, trade name, advertising slogan, specification, or confidential information, but is solely intended to provide a revocable license to use the specification material for the contemplated use.

Project:
11/25/03
Section E - 1

