Part 1: Engineer’s Specification

## Overview

It is the intention of ________________ (Facility Name) to install a single, facility-wide, Nurse/Patient Communications Network that offers virtually unrestricted flexibility in assignment of annunciation for system patient stations, peripherals, and nurse consoles. This system shall be a single network capable of up to: 20 Central Equipment modules, supporting up to 40 Master Station consoles, up to 640 Patient Station annunciation points, and over 7,040 Peripheral Stations.

Conventional nurse call systems that utilize discrete area control units, limited Personal Computer operating systems, or local controller/switchers shall not be accepted.

All system components shall be listed to the UL®1069 standard by a qualified NRTL. Systems having a “core” system only listed shall not be acceptable.

## Section Includes

A Nurse/Patient digital communications network is to be installed initially on floor(s) __________ and then expanded in the future to the remaining units.

## References

- Underwriter’s Laboratories ANSI/UL®1069

- Canadian Standards Association C22.2, No. 205

- National Electrical Code, ANSI/NFPA 70

- Health Care Facilities, ANSI/NFPA 99

- U.S. Dept. of Labor / Occupational Safety and Health Administration

- State Hospital Code / Joint Commission of Hospitals—Nurses Call Requirements

- ISO 9001 Quality Standard

## Qualifications

- Authorized Distributor for product supplied. Authorized Distributor Letter from manufacturer required upon request of specifying authority.

- Applicable state licenses.

- A current certificate of successful completion of manufacturer’s installation/training school for installing technicians of the equipment being proposed.

  - Training recertification to occur every 3 years

## System Description

System hardware shall be manufactured by an ISO 9001 certified company. Equipment manufactured by companies that are not ISO 9001 shall not be acceptable.

The hardware shall consist of the CENet nurse/patient communications network, comprised of:

NC450 series Central Equipment (CE),

Master Station Consoles (two configurations):

NC402TS (LCD Touch Screen), or

NC401VFD (4×40 Vacuum Fluorescent Display).

Patient stations, bath pull-cord stations, bed side-rail interfaces, code call stations, corridor dome lights, pillow speaker/entertainment cords, call cords, emergency pull-cord stations, emergency push button stations, lamp control modules, presence stations, radio page interface, remote jacks, computer interface, printer interface, and wiring shall be supplied as shown on the drawings.

All necessary equipment required to meet the intent of these specifications, whether or not enumerated within these specifications, shall be supplied and installed to provide a complete and operating nurse/patient communications network.

The system shall provide for full-duplex audio between master stations in the handset mode, and remote (patient, staff and duty) audio stations. Half-duplex audio shall be selectable by individual remote audio station for areas with abnormal acoustical requirements.

System firmware shall be the product of a reputable firmware manufacturer with a proven history of product reliability and sole control over all source code. Manufacturer shall provide, free of charge, product firmware/software upgrades for a period of five years from date of installation for any product feature enhancements. System firmware upgrades shall not require any exchange of parts. Any supplier whose equipment requires the exchange of parts for firmware upgrades shall not be acceptable.

Equipment manufacture and testing shall be executed by an ISO 9001 certified company. Manufacturing testing shall utilize applicable custom fixtures to assure the highest quality production.

The Nurse Call System provider shall provide a Fail-Safe Network topology as described below:

- The Network shall be of a closed, proprietary nature, except where routing technology is employed to provide limited, controlled access to external ancillary systems. The Network bandwidth shall be automatically monitored and the traffic shall be automatically controlled to ensure the specified performance.

- The Central Equipment must be able to function as a Network using a bus topology where a Central Hub is deemed undesirable. A single failure of the Central Equipment shall not affect the functioning of the system.

- The Network shall be electrically supervised and include verification protocols for all critical data transmissions.

- All components on the Network shall be provided with battery backup power in the event of the loss of facility-supplied AC power.

- All system data required for the Network’s operation and that of its components as relates to the primary nurse call functions shall reside in non-volatile solid-state memory. In the event of catastrophic failure due to adverse environmental conditions, the system shall use the most recent data to recover to normal operation.

## Submittals

The system described by this specification is the Tek-CARE®400 system using PM453 modules, manufactured by TekTone® Sound and Signal Mfg., Inc. The Tek-CARE®400 system using PM453 modules meets all requirements outlined in this specification. The Tek-CARE®400 system using PM453 modules should be considered the basis for all submitted bids.

Any supplying contractor proposing equipment which is not the base standard for this specification must provide full submittals at the time of bid. This option shall be exercised at the discretion of the owner/specifying authority.

Prior to commencement of work, the supplying contractor shall submit six (6) complete submittal sets. These sets are to be submitted in three-ring binders, continuous spiral binders, or plastic bindings that allow the booklets to lie flat while open. Each booklet shall consist of the following:

Page 1:	Name of supplying contractor and project name

Page 2:	In the following order, a listing of: component quantities, equipment manufacturer, model number, and description of each component being supplied. If equipment being supplied is not the specified equipment manufacturer’s model, alongside the submitted model number and description, list the specification paragraph that corresponds to the equivalent specified model. Failure to provide this information shall result in the rejection of submittals.

Page 3:	Recently dated (within one year from submittal date) support letter from manufacturer stating that the supplying contractor is an Authorized Distributor of the product being supplied.

Page 4:	A statement of warranty policy from manufacturer.

Page 5:	A copy of the installing technician(s) certificate of completion from the manufacturer’s training school (within three years from submittal date) for the equipment being proposed.

Page 6:	A written statement by contractor of how and when they will complete In-Service Training, including the minimum number of hours being provided per system, procedures they will follow, what training aids are provided (technical and user manuals, data sheets, etc.) and how contractor will conduct training.

Page 7:	A written statement from contractor of: (a) exactly how the contractor will test installed equipment and wiring, and (b) exactly how all the tests recommended by manufacturer will be performed by the contractor, prior to commissioning of system.

Page 8:	Provide a written copy of the manufacturer’s list of all of the recommended spare parts to maintain all systems specified after the warranty period. Provide the purchase price and turnaround cost (i.e., Facility’s) associated with each item. List separately the cost of an annual maintenance contract. Show the hourly, purchased labor rates for both regular and emergency service. State any additional charges that may accompany labor charges (such as, but not limited to, travel charges, lodging, etc.).

Pages 9+:	One catalog sheet per product of equipment listed on page 2, in the exact order as listed on page 2. Each catalog sheet shall describe mechanical, electrical and functional equipment specifications. Photocopy duplications of the manufacturer’s original equipment catalog sheets will be allowed as long as they provide adequate clarity of all printed words, graphics, pictures, illustrations and other information material to the evaluation of the submittal. Submittals that are not of adequate clarity or content may be rejected and re-submission may not be allowed.

Last Page(s) or Separate: Provide all inter-equipment wiring diagrams and drawings necessary to install the equipment being supplied. These drawings will show, in detail, all wiring types by wire gauge, conductors and wire manufacturer. These drawings must be updated prior to completion of any work to reflect changes that may have been made during actual installation.

In the event the specifying authority decides to reject the submittals of a supplying contractor, the specifying authority may ask the contractor to re-submit if the discrepancies are minor. Otherwise, rejection of submittals means the specified product must be supplied.

## Project Site Visit

It is the responsibility of all prospective contractors to make an adequate inspection of the project site or review of project plans. A mandatory site visit or meeting to review project plans is scheduled for the following date(s) ________________. Any contractor not registered as having attended the mandatory site visit tour or the project plans review meeting will be disqualified and any bid proposal will automatically be rejected.

## Demonstrations

It may be necessary to utilize demonstration equipment to test the functional operation of the contractor’s submitted equipment. Contractor will be notified of any demonstration dates and times. If such demonstrations are utilized, it will be the sole judgment of the owner and/or owner’s designated specifying authority to decide whether a contractor/manufacturer meets or exceeds the specification.

All demonstrated equipment must be of a standard single manufacturer and meet the same required testing and conditions that are applicable to the manufactured equipment. Custom or modified equipment that is not of standard, current manufacture cannot be demonstrated.

If necessary, owner and/or specifying authority may visit the manufacturer’s facility to view functioning equipment or demonstrations and witness equipment manufacturing techniques and/or testing procedures.

## Samples

The owner/specifying authority reserves the right to request one each, samples of terminal (station) equipment for the purpose of coordinating colors, aesthetics, trim-plate sizing, etc. These samples would be supplied at no cost to the owner after bid is awarded.

## Scheduling

It is the responsibility of the general contractor to coordinate all work with the other trades for scheduling, rough-in, and finishing all work specified. The owner will not be liable for any additional costs due to missed dates or poor coordination of the supplying contractor with other trades.

## Warranty

The warranty shall include all necessary labor and equipment to maintain the system(s) in full operation for a period of one year from the date of acceptance. Equipment only, manufactured by the manufacturer shall be warranted for five years; provided the installation is performed by factory-certified technicians, and an inspection of the installation is done by a person(s) designated by the factory. The contractor shall maintain a service department, necessary spare parts, telephone answering services, and call dispatching required to implement the service standard stated below as part of this contract. After the acceptance of the system(s), service shall be provided on the following basis:

Emergency Service—Provided 24 hours a day. When a total or catastrophic failure of equipment is reported to contractor, within 2 hours of notification, a service person will be on site. (An example of a catastrophic failure would be a central equipment failure or a nurse master station failure.)

Routine Service—Provided within 4 business hours (9 a.m. to 5 p.m., Monday through Friday, excluding holidays) of notification. When a minor failure of equipment is reported to the contractor on a Saturday or Sunday, a service person will be on site within 4 hours of notification. (An example of a minor failure includes equipment such as patient stations, corridor lights, pull-cord stations, etc., which normally affect only one patient or patient room.)

## Maintenance

The contractor shall:

- Provide necessary spare parts, noted on Page 8 of submittal (see Section 1.5), after commissioning of system(s) and before final payment.

- Provide sponsorship for at least one person designated by the owner to attend a service school held by the equipment manufacturer. Transportation to this school, meals, and lodging will be borne by the owner. The equipment manufacturer shall provide school free of charge.

- The owner may choose to have the supplying contractor maintain the system(s). The level of service provided during the maintenance contract period would be the same as the warranty period for routine and emergency service. All labor and equipment costs shall be covered under this contract. Supplying contractor must state exact billing amounts, billing periods and all costs associated with this maintenance agreement and list any items that would not be covered under the service/maintenance agreement.

## Part 2: Products

## Manufacturers

The products specified shall be new and of the standard manufacture of a single, reputable, ISO 9001 certified manufacturer. As a reference of standard and quality, functionality and operation, it is the request of the owner that bids be based only on equipment manufactured by TekTone® Sound & Signal Mfg., Inc., Franklin, NC 28734.

## Removal of Existing Product

The contractor shall remove all existing product and deliver to the owner, or at the direction of the owner, properly dispose of same.

A.	The owner will vacate a nursing unit at a time, making it available on a timetable for the installation of the new equipment.

–or–

B.	The owner will continue to occupy the nursing units where equipment will be replaced. Supplying contractor will need to coordinate work with nursing administration for each nursing unit to obtain a minimum group of four rooms at a time for replacement of equipment. The existing nurse call equipment or a temporary wireless call system must be maintained and operational during this replacement period, except for the four rooms being renovated.

## Network Wiring

There shall be five (5) types of wiring within the system.

- Commons Bus shall be CAT5 or better cable and 6P6C modular connectors. One pair in the CAT5 cable shall remain unused. Cables shall be wired straight-through. Commons bus shall not exceed 1000’ of cabling.

- Station Power Circuit shall be #16-18AWG pair plus an earth ground. TekTone® recommends using a 3-conductor cable for power and ground. Station Power Circuit shall not exceed 1000’ of cabling.

- Peripheral Bus shall be CAT5 cable and 8P8C modular connectors wired using the T586B standard.

- Dome Light Bus shall be CAT5 cable and 8P8C modular connectors, wired using the T586B standard. The dome light bus shall also include power on an #16-18AWG pair plus and earth ground. TekTone® recommends using a 3-conductor cable for power and ground.

- Master Station Wiring

- Power, 1 pair 14-16AWG. Manufacturer shall provide chart for installer to determine lesser wire gauge requirements for shorter cable distances.

- #18–#24 AWG, 4 twisted pairs (CAT5 or better is acceptable) and 8P8C modular connectors.

- NC402TS utilizes a second CAT5 cable to Tek-CARE® Network to enhance database access and software update distribution.

- Central Equipment Network (CENet) Wiring shall meet standard Ethernet requirements between CEs and switches as needed.

## Central Equipment (CE)

The contractor shall furnish, as shown on the plans, an NC450 Central Equipment to furnish each nursing unit with the listed components as needed (numbers in parenthesis reflect maximums within an NC450). The NC450 provides connections for up to 7 modules from the list below. The system shall support additional modules required at a specific location by adding NC450 devices as needed:

1—NC450 CE Housing with dimensions no greater than 19.5" wide × 24.25" tall × 8" deep, with PM450 Backplane Module mounted internally.

1—PM451 Hub Control Module with system network data interface port.

1(5)—PM453 Master & Station Module.

0(5)—PM455 P5 Master & Station Module.

0(5) PM456 NC300II Head End Module.

2(6)—PK450 +28 VDC Primary Power Supply.

2(6)—BA450K Battery Backup Kits (provides supported module with 10 minutes of backup power). Provide an auxiliary source of power in the event of the failure of the primary electrical power source. See ANSI/NFPA 99, Health Care Facilities.

Or

1—NC450MICRO CE (one PM453 module per NC450MICRO)

Central power shall be monitored at each PM453, PM455 and PM456 Master & Station Module. In the event that a switch is made to battery power, “Battery Backup Active” (default text) shall be displayed at each master station after the voltage reaches approximately 24VDC.

It shall be possible for each PM453, PM455 and PM456 Master & Station Module to act as a standalone controller for its local stations and master stations in the event of a loss of network communications. Should such a loss occur, the PM453, PM455 or PM456 Master & Station Module shall issue network faults to local master stations.

## Nurse Master Stations (Consoles)

- The contractor shall furnish, as shown on plans, optional nurse console(s) Model NC402TS, a compact Nurse Master Station with a backlit color touch-screen LCD VGA Display (10.4" viewing area). It shall be a self-contained unit, which shall not occupy more than 117 square inches of desk space. Model NC402TS shall be capable of the following functions:

- The touch screen of the NC402TS shall be the primary interface used to control, access, and view calls displayed on the touch screen.

- The NC402TS shall contain a membrane keypad with redundant functionality used to control, access and program system features and calls displayed on the touch screen.

- The touch screen displays current calls and other data.

- Full English display with user prompts.

- Ability to display data on the current locations of staff.

- Water-resistant retractable keyboard with rubber overlay and redundant membrane switch functionality to allow for easier entry of data.

- The NC402TS shall support connectivity to the Tek-CARE® network via Ethernet.

- Connectivity for a standard USB keyboard and/or mouse shall be supported via USB ports on the rear of the NC402TS master station.

- The NC402TS shall display up to 20 incoming calls with the ability to select a call and/or scroll to any active call using the touch screen, keyboard, membrane switch or USB mouse. Each call shall have an individual elapsed timer which increments until the call is reset. Alternatively, calls and staff location may be displayed on station icons arranged by user preference into lockable positions on the touch screen.

- Master Station is powered from the NC450 Central Equipment. Separately powered Nurse Master Stations requiring access to local AC power, separate console power supplies, and/or requiring local UPS backup power units are not acceptable.

- The Master Station shall be able to receive and display any or all calls placed in the system, including simultaneous call types from the same room. Calls shall be sorted and displayed first by call priority and then by the chronological time in which the calls were placed.

- There shall be at least 256 possible unique user-definable call types.

- Choice of Push-to-Talk or private conversation using the handset.

- Automatic answer of highest priority call or selective answer of any displayed call.

- Set/Review up to four levels of service required—STAT Assist, Staff L1 (Green), Staff L2 (Amber) & Staff L3 (Yellow).

- Day/Night call-in tone-level control.

- Audio—All Page, Zone Page and Staff Page operation.

- Optional tone silence by user definable call types. Tones are regenerated whenever a new incoming call is received.

- Tone silence defaulted to “Routine” calls.

- Software shall be able to optionally defeat Tone Silence feature.

- Optional muted tone for call waiting indication during audio communication.

- Ability to block all loudspeaker paging to facilitate a low-noise patient environment. Password protection can be enabled to allow only authorized access to audio paging.

- Ability to swing an individual room using convenient per-station zoning.

- Ability to zone capture an individual nursing unit, selected units, or all units in facility by using custom-defined zones, per-master zoning and per-station zoning.

- Ability to do day/night transfer between consoles by selecting or dialing a master and initiating a Master Forward.

- Direct messaging (canned/custom) to pocket pager(s).

- Set/review patient privacy.

- Locate up to three levels of staff with remote cancel of manual staff registration: Staff L1 (Green), Staff L2 (Amber) and Staff L3 (Yellow).

- Room monitoring with room number display.

- Continuously supervised with faults displayed as real time error messages.

- ESD protected to 8kV per UL®1069.

- Consoles may be located anywhere within facility nurse/patient communications network.

- Top half (180-degree) spill-proof design to work in conjunction with spill-proof membrane switch to prevent liquid spill damage to master station.

- Pleasant sounding call tones.

- The contractor shall furnish, as shown on the plans, standard nurse console(s) NC401VFD compact Master Station with an integral 4×40-character Vacuum Fluorescent Display (VFD) and spill-proof membrane switch. It shall be a self-contained unit, which shall not occupy more than 117 square inches of desk space with the following specifications:

- Master Station is powered from the NC450 Central Equipment. Separately powered Nurse Master Stations requiring access to local AC power, separate console power supplies, and/or requiring local UPS backup power units are not acceptable.

- 4×40, 160-English-character backlit VFD display; Alphanumeric Room Display with Bed and Call Priority.

- Display up to 4 incoming calls (with the ability to scroll to any active call), each with an individual elapsed timer that increments until the call is reset.

- The Master Station shall be able to receive and display any or all calls placed in the system, including simultaneous call types from the same room. Calls shall be sorted and displayed first by call priority and then by the chronological time in which the calls were placed.

- There shall be at least 256 possible unique user-definable call types.

- Choice of Push-to-Talk or private conversation using the handset.

- 20 programmable buttons with LEDs on Membrane Switch for representing stations, masters, zones and or system features.

- Automatic answer of highest priority call or selective answer of any displayed call.

- Set/Review up to four levels of service required—STAT Assist, Staff L1 (Green), Staff L2 (Amber) & Staff L3 (Yellow).

- Day/Night call-in tone-level control.

- Audio—All Page, Zone Page and Staff Page operation.

- Optional tone silence by user definable call types. Silenced tones are regenerated whenever a new incoming call is received.

- Tone silence defaulted to “Routine” calls.

- Software shall be able to optionally defeat Tone Silence feature.

- Optional muted tone for call waiting indication during audio communication.

- Ability to block all loudspeaker paging to facilitate a low-noise patient environment. Password protection can be enabled to allow only authorized access to audio paging.

- Ability to swing an individual room using convenient per-station zoning.

- Ability to zone capture an individual nursing unit, selected units, or all units in facility by using custom-defined zones, per-master zoning and per-station zoning.

- Ability to do day/night transfer between consoles by selecting or dialing a master and initiating a Master Forward.

- Direct messaging (canned/custom) to pocket pager(s).

- Set/review patient privacy.

- Locate up to three levels of staff with remote cancel of manual staff registration: Staff L1 (Green), Staff L2 (Amber) and Staff L3 (Yellow).

- Room monitoring with room number display.

- Continuously supervised with self-diagnosing error messages.

- ESD protected to 8kV per UL®1069.

- Consoles may be located anywhere within facility nurse/patient communications network.

- Spill-proof membrane switch to prevent liquid spill damage to master station.

- Pleasant sounding call tones.

- Master Station firmware operates on TekTone’s operating system. Master Stations which utilize an outside supplier’s operating system where software failure (“lock-up”) may occur due to inconsistencies and incompatibilities between operating system and equipment supplier’s software, rather than operating in a firmware environment, or Master Stations requiring mechanical hard disk drives, which are prone to mechanical failure, shall not be accepted. All required software/firmware shall be supplied by TekTone®.

- Programming of Master Stations shall be done via the integrated Master Station user interface or from a centrally connected PC-based configuration tool.

- Staff-to-patient RF paging assignments shall be easily programmable from the NC402TS console at the beginning of each shift by selecting the staff member’s name and changing their personal assignment screen. The staff-to-patient assignments are established directly using the patient’s room and bed. Additionally, the staff may receive all 256 call types, or only a selected subset of call types. Master Stations that do not allow staff to program patient assignments by staff name are not acceptable.

- It shall be possible to remove and/or replace any console(s) while the system is operational without the loss of any calls, damage to any system components, or reprogramming of console attributes.

## Bedside Patient Stations

The contractor shall provide single bed Model IR411 or dual bed Model IR412 as shown on plans. Each IR411 single or IR412 dual bedside patient station shall be capable of the following:

- At least 256 completely custom configurable call types. Configurability shall extend to call labels/priorities/levels/tone-dome light annunciations.

- At least 38 default call types (including associated peripherals).

- Bath Call

- Bath Fault

- Bed Out

- Code Call

- Code Fault

- Communications Fault (“Comm. Fault”)

- Cord Out, ¼" / DIN left side (IR412 only)

- Cord Out, ¼" / DIN right side

- Bed Out, Hill-Rom®

- Emergency Call

- Emergency Fault

- Emergency2 Call

- Emergency2 Fault

- Auxiliary Input 1A

- Auxiliary Input 1B

- Auxiliary 1 Fault (supervises the module, not the device it is connected to)

- Auxiliary Input 2A

- Auxiliary Input 2B

- Auxiliary 2 Fault (supervises the module, not the device it is connected to)

- Auxiliary Input 3A

- Auxiliary Input 3B

- Auxiliary 3 Fault (supervises the module, not the device it is connected to)

- Lamp Fault

- Monitor Call

- Personal Attention Call

- Priority Call

- Routine Call

- Staff Emergency

- Staff L1 Presence

- Staff L2 Presence

- Staff L3 Presence

- Staff L1 Request

- Staff L2 Request

- Staff L3 Request

- Remote JK400 Cord Out Left DIN Jack. (IR412 only)

- Remote JK400 Cord Out Left ¼" Jack. (IR412 only)

- Remote JK400 Cord Out Right ¼" Jack

- Remote JK400 Cord Out Right DIN Jack

- Programmable 24-character patient name or label, plus an 8-character architectural room name that display at master station with other relevant call information.

- Four levels of staff service.

- Full duplex audio with the master station from the handset.

- Include single (IR411) or dual (IR412) DIN jacks for entertainment/pillow speaker and ¼" jack(s) for call cords or dry contact inputs.

- Dummy plugs shall not be required to prevent calls from empty patient front panel station connectors (for ¼" or DIN receptacles). Systems that require dummy plugs to prevent calls are not acceptable. Patient stations must provide “cord-out” detection when call cords are used.

- Include single (IR412) or dual (IR412) Hill-Rom®/Stryker bed Interface (includes bed exit alarms), Intercom, Lamp Module (PM421/PM422). Module interface shall be polarized for proper module insertion.

- Include Dome Lamp (two-, and four-bulb versions), designed to connect directly to patient station via current-limited station outputs. (LI482LED, LI484LED)

- Common call-reset button for all Routine, Priority and Upgraded calls placed from station and one call-placed LED per bed.

- Protected to 30,000 volts Electrostatic Discharge (ESD) with test verification performed by Underwriters Laboratories, Inc. Compliance with 30,000 volts ESD protection of non-specified alternative system requires performance test at Underwriters Laboratories, Inc. at the discretion of the owner/specifying authority.

- Continuous supervision for station power and data lines as well as data communications.

- Plug-in field wiring connections.

- Ability to service exchange station “hot” (i.e., without removing system power or powering down local Central Equipment).

- Ability to program on a per station basis, each bed and entertainment/call cord receptacle and/or bused peripheral input to a custom call type.

- Patient station addressing shall be accomplished by simple dip switch addressing. Methods that do not provide for simple addressing means shall not be acceptable.

- Provide connections and functionality for a peripheral bus supporting the following peripherals:

- Bath Station*

- Code Call Station

- Emergency Station*

- Emergency2 Station*

- 3-position staff presence with staff emergency button

- Up to six (6) isolated interface inputs to local equipment alarm contacts (i.e., ventilator, IV drip, etc.) to notify Master of local alarm condition in patient room. Each input shall default to 6 call types—programmable to any one of 256 custom call types. This shall be supported in a modular format allowing users to add isolated inputs in groups of two (2).

Peripherals marked with an asterisk (*) shall be supported in multiple quantities by patient station with a maximum of three duplicate devices using a substitution method. (Example: Bath, Code, Emergency and Emergency2 peripherals could be changed to Code, Emergency, and two Bath devices.)

Note: All peripherals shall be supervised for connection to the controlling station; the Code station shall be fully supervised.

- Patient station shall fit within a RACO 692 (or equivalent) housing when PM421 or PM422 module is not present. When the module is used, the assembly shall fit within the RACO 697 (or equivalent) housing.

-or-

The contractor shall provide single bed Model IR411EPB or dual bed Model IR412EPB as shown on plans. Each IR411EPB single or IR412EPB dual bedside patient station shall perform the same functionality as the IR411 single or IR412 dual bedside patient station with the following feature difference: In lieu of the ¼" jack, there shall be a momentary button on each Patient Station that defaults to a Staff Emergency call. The momentary button(s) shall be programmable to a custom call type.

## Multipurpose Station

The contractor shall provide as required, a universal interface Model IR400 Multipurpose Station for standalone peripheral stations. Where shown on plans, this module may be used to drive corridor paging amplifier(s) via a grounded, shielded transformer (to prevent leakage current from being impressed on the isolated circuits of the nurse call system). Alternately, the IR400 may be connected directly to a 25-volt, ½-watt speaker for paging with listen-back. Always use grounded, shielded transformers to prevent current leakage to the nurse call system’s isolated circuits. This unit shall also be capable of providing auxiliary alarm device contacts inputs for nurse console notification of a local alarm(s) where necessary. It shall also provide the same peripheral inputs as patient stations.

## Staff Station

The contractor shall provide as required a Staff Station Model IR410. Each IR410 or staff station shall have the same functionality as the IR411 single bedside patient station with the following feature differences: In lieu of the call jack(s), a call button shall place a call to the assigned Master Station; optional programming shall allow the default staff-call type to be changed to any valid patient station call type; and in addition, the station features an integrated emergency call button.

## Duty Stations

The contractor shall provide as required a Duty Station Model IR415. Each IR415 duty station shall have the same functionality as the IR411 single bedside patient station with the following feature differences: In lieu of the ¼” and DIN jack, a call button shall place a call to the assigned Master Station; the unit shall provide remote annunciation of assigned bedside patient stations and peripherals via call placed LED, emergency LED and call tones; it shall drive four (4) lamp corridor light versions (in a zone lamp mode); it shall be tested for protection to the UL® 1069 standard for Electrostatic Discharge (ESD); and it does not support Hill-Rom®/Stryker/Sunrise Medical CCG Interfaces or Pillow Speaker Modules.

## Peripheral Stations

The contractor shall provide, as shown on plans, peripheral stations associated with Patient, Multipurpose, Staff or Duty stations wired via an RJ connector using peripheral bus scheme. Peripheral devices that do not support a bus scheme or require individual direct wiring connections to station are not acceptable. Individual peripheral stations shall be:

- Pull-cord peripheral Emergency Shower Switch Model SF437 shall be water-resistant (when installed according to manufacturer’s instructions with included gasket kit). The station shall have the ability for user defined customized call labels/priorities/levels/tone-dome light annunciations, corresponding to up to 256 software call types available within the system.

- Addressable pull-cord Emergency Shower Station for standalone areas such as pool, tub or bathrooms shall be Model SF483. The station shall be built on a single-gang ABS plastic faceplate and shall use magnetic reed switches for contact closure, O-rings for screws and a gasket for mounting against the wall or tile shall provide for waterproof installation. A front replaceable pull cord and large red pull-down actuator shall allow the station to be used with or without the pull cord. It shall be possible to select user-defined customized call labels/priorities/levels/tone and dome light annunciation. The station shall be connected directly to the station commons bus wiring of the system. Standalone emergency pull-cord switches that require an additional module for interface shall not be acceptable.

- Single call latching button Emergency Switch, Model SF439, shall have a mushroom switch. The latching button shall have an associated call assurance LED. The switch allows the staff to locally cancel the call by twisting the button. The station shall have the ability for user defined customized call labels/priorities/levels/tone-dome light annunciations, corresponding to up to 256 software call types available within the system.

- The Model SF440 Emergency Switch is intended for locations where emergency signaling is required. Activation is achieved by pulling the pull-cord or by sliding the “Emergency” slide switch on the single gang panel. The station shall have the ability for user defined customized call labels/priorities/levels/ tone-dome light annunciations, corresponding to up to 256 software call types available within the system. Each station shall have an associated call assurance LED. Sliding the “Emergency” switch to the reset position allows the staff to locally cancel the call.

- Addressable pull-cord Emergency Station for standalone areas such as public bathrooms shall be Model SF481. The station shall be built on a single-gang ABS plastic faceplate with front replaceable pull cord and large red pull-down actuator so it can be used with or without the pull cord. It shall be possible to select user-defined customized call labels/priorities/levels/tone and dome light annunciation. The station shall be connected directly to the station commons bus wiring of the system. Emergency pull-cord switches that require an additional module for interface shall not be acceptable.

- The Model SF441 Code Call Switch peripheral station shall have the ability to place a network-wide code blue annunciation at specific master stations, and/or centralized master station. Pocket page code team activation shall be automatic. In addition to power and data supervision, the IR400-series unit shall have the ability to directly supervise the associated SF441 Code Call Switch and a supervised type corridor light for open, shorts and missing bulbs.

- Addressable Code Call Station for stand-alone areas such as clinic or procedure rooms shall be Model SF482. The station shall be built on a single-gang ABS plastic faceplate with large blue pull-down actuator. It shall be possible to select user-defined customized call labels/priorities/levels/tone and dome light annunciation. The station shall be connected directly to the station commons bus wiring of the system. Stand-alone Code Call switches that require an additional module for interface shall not be acceptable.

- Presence Station Model SF450/4 shall have four (4) color-coded switches associated with user defined staff levels [Staff Assist (Red), Staff L1 (Green), Staff L2 (Amber) & Staff L3 (Yellow]. When a staff member presses the assigned switch, the associated LED and corridor staff lamp shall light. An existing patient call or service requirement shall be automatically canceled at time of registration and the staff may be located from any assigned nurse console. When the associated nurse console is placed in “Staff Follow” all call tones from this assigned console shall also sound over the staff register peripheral’s associated Patient Station.

## Patient Entertainment Speaker/Call Cords

The contractor shall:

A.	Provide, one per bed, an entertainment pillow speaker Model _______ that includes the following functions:

Model SF41ZDL: Nurse call button designated by universal nurse symbol, 8' cord with heavy duty DIN plug, TV channel up/down, volume control up/down, TV power, mute, closed caption, sleep, room light, and read light controls. The pillow speaker shall be electrostatic discharge (ESD) protected to over 12,000 volts, and shall be oxygen safe.

Model SF41ZKL—same as SF41ZDL, plus: keypad for direct channel access instead of channel up/down.

B.	Provide, one per bed, a DIN-plug strain relief cable model SF401EX.

C.	Provide, one per 10 beds or as required, a Model SF401A single-pendant-type oxygen-safe/waterproof call cord complete with 7' cord, DIN plug and sheet clip.

D.	Provide, one per 10 beds or as required, a Model SF401G single oxygen/geriatric pressure-ball cord set utilizing a pneumatically controlled switch for use in oxygen-enriched environment, complete with 6' cord, DIN plug and sheet clip.

E.	Systems requiring special adapters to convert to DIN plugs, or that use non-standard (non-DIN) cord sets, shall not be acceptable.

## Corridor Light Sets

The contractor shall:

Provide corridor lights as shown on plans. Lights with the “SUP” suffix shall provide supervision of lamp bulbs. Lights with the “Z” suffix shall provide zone light functions (echoing of light information from assigned zones) and supervision of lamp bulbs. Lights with the “LED” suffix shall provide LEDs instead of incandescent bulbs. All lights shall be interchangeable.

- Corridor lights with LEDs shall be available with the following number of lamps:

- Four (4) LED corridor light Model LI482LED, with 2 white and 2 red LEDs.

- Four (4) LED corridor light Model LI484LED-1, with white, red, green & amber LEDs.

- Four (4) LED standalone zone light Model LI484ZLED, with white, red, green & amber LEDs.

It shall be possible to use the system’s programming features to reassign lamp assignments. Defaults are as follows (listed alphabetically):

- Bath Call—flashing Red

- Bed Out—flashing White

- Code Call—alternating Blue and Red

- Code Fault—None

- Communications Fault (Comm. Fault)—None

- Cord Out (any)—flashing White

- Emergency Call—flashing Red

- Emergency2 Call—fast flashing Red

- Isolated Inputs Six (6)—programmable

- Lamp Fault—none

- Monitor Call—flashing White

- Personal Attention Call—steady White

- Priority Call—flashing White

- Routine Call—steady White

- Staff Emergency—flashing Red

- Staff L1 Presence (Level 1)—steady Green

- Staff L2 Presence (Level 2)—steady Amber

- Staff L3 Presence (Level 3)—steady Yellow

- Staff L1 Service Request (Level 1)—flashing Green

- Staff L2 Service Request (Level 2)—flashing Amber

- Staff L3 Service Request (Level 3)—flashing Yellow

It shall be possible to program a custom light scheme to provide any color combination (White, Red, Blue, Yellow, Amber and Green) and steady or flashing for each of the 256 available call priorities. System shall provide indication to user when programming a call type and/or staff indicator conflict. If a conflict is permitted, system shall display call types over staff indications.

## Data Interface Port on Hub Control Module

The contractor shall furnish as necessary, a Hub Control Module with data interface port, Model PM451, to be used with the NC450 and to provide the following features:

- System programming via network port of all Network Central Equipment, Masters, Stations, assignments and entire network operation from a single interface point.

- Upload all network component firmware to flash memory on site.

- Network administration of pocket page setup.

- Laptop interface for diagnostics of all network components.

- A hardware interface for the TekTone® Tek-BRIDGE®PC Model NC474.

- Systems that require local programming of individual network components/subsystems at multiple physical locations, or that do not allow for remote upload of component firmware shall not be accepted.

- Systems that require a separate module for data interface shall not be acceptable.

## Integration Server

The contractor shall:

The contractor shall furnish an NC474 Tek-BRIDGE® Server with Tek-CARE® software as required.

The Tek-BRIDGE® server shall provide the platform to install and run the ancillary software modules for Tek-CARE® healthcare communications systems and the Tek-ALERT® alert integration system.

It shall also supply ports for the Tek-CARE® network; a local area network (LAN) or the facility’s LAN; the radio pocket paging system’s paging transmitter; Tek-CARE®300 II central equipment; the NC502 Central Monitoring Auto Dialer, NC510ES EchoStream wireless Receiver, and NC510 FA Receiver, Tek-CARE® wander management system and license network integrations for AppleTV® and mobile iOS® products.

The Tek-BRIDGE® includes the LS450 Config Tool, which is used to set up and customize Tek-CARE® and Tek-ALERT® systems and their add-on components, and to save the system configuration.

The integration server shall have the following minimum features.

- Solid state hard drive. Servers with mechanical hard disk drives shall not be accepted.

- One (1) VGA monitor port.

- Three (3) Ethernet ports.

- One (1) PS/2 keyboard connection and one (1) PS/2 mouse connection.

- Six (6) DB9 serial connections.

- Four (4) USB ports.

- 3.5mm Speaker jack, 3.5mm Mic jack, and 3.5mm Line In jack.

- Internal speaker for annunciating call tones.

- Pleasant sounding call tones.

- Preloaded with appropriate configuration software for all connected systems.

- Provided as a complete unit, with monitor, keyboard, and mouse included.

- The integration server shall not require an IP connection for backup or software upgrades. Systems which require an IP connection to manufacturer servers shall not be accepted.

- The integration server shall accept software licenses to enable connections to various native and foreign systems.

- The integration server shall accept a license to communicate with both current and legacy TekTone® nurse call systems as required. The license shall be TekTone® LS301.

## Event Monitor Software

The integration server shall be provided with proprietary event monitor software for displaying and routing calls received by the Nurse Call System. The event monitor software running on the integration server shall display all calls from connected analog Nurse Call Systems as well as Nurse Call System faults. The event monitor software shall be TekTone® LSMON-series. The system shall support up to ten (10) event monitors if so licensed.

Any event that is annunciated by the system shall be automatically logged in the reporting database. The database shall record all system activity for review at a later time using the optional Reporting software. The database shall record all information about the event, including response time and the time and date of the event. The database shall be operationally transparent to the Nurse Call System. No action by the user shall be necessary to record system events. The database shall not require any maintenance or periodic cleanup by the user.

The event monitor software shall also enable users to view information about the system, including patients, staff, staff groups, staff assignments, and scheduled messaging.

The Nurse Call System shall support an optional remote event monitor that may be installed on any networked facility personal computer running Windows XP or higher. The remote event monitor requires that the integration server be connected via LAN and properly licensed.

## Configuration Software

The integration server shall be supplied with configuration software for system setup and programming. The configuration software shall enable programming of all licensed system features.

## Reporting System

A reporting system that is operationally transparent to the Nurse Call System shall be provided with the system. The reporting database shall automatically log system events and store them for retrieval at a later date.

The reporting system shall be accessible from the event monitor software running on the integration server or remote event monitors running on networked facility PCs. Report creation shall be simple and completely customizable. Reports shall be viewable in a web browser as an HTML file or exported as a CSV file. The reporting system shall have quick links for creating commonly requested report ranges (last day, last week, etc.) as well as offering fully customizable report creation.

Using the optional LS453 Email Output software, the reporting system shall be capable of generating reports and distributing them automatically via email.

The reporting system shall be available as an optional feature and shall be TekTone® LS610 series.

## Pocket Paging

The contractor shall;

Provide a pocket paging transmitter and alphanumeric pagers, or LED Messaging Sign. The transmitter shall be capable of being connected to the Nurse Call System via the Tek-BRIDGE® Server, and shall be operationally transparent to the Nurse Call System. The pocket paging interface shall allow COMP 2, TAP and Scope protocols. The pocket paging equipment shall support the following features:

- The Nurse Call System shall accommodate up to 256 unique pager cap codes. Each pager cap code may be given a name describing it and/or the staff member who carries it. Each pager may belong to multiple pager groups, of which the system shall allow up to 256.

- The system shall allow four types of paging assignment:

- From a station to a pager

- From a station to a pager group

- From a zone to a pager

- From a zone to a pager group

- Each assignment shall have the following properties:

- Level—whether the assignment is primary, secondary, or tertiary

- Call filter—the call types to which the assignment applies

- Page resets—whether the assignment applies to resets

- Page faults—whether the assignment applies to faults

- Paging assignments shall be fully configurable from the nurse call master station, remote event monitors and from the Tek-BRIDGE® Server running the Config Tool programming software.

The system shall allow two types of paging: manual and automatic. Automatic paging may be turned off or on; when on, it shall operate concurrently with manual paging.

Manual paging shall be executed either from the integration server, remote event monitor or a nurse call master station. At any time, the attendant may page by pager, pager group, station or zone. For convenience, system paging shall also be directly supported. Paging by station or zone shall occur through that station’s or zone’s paging assignments. For a station page, the attendant may select the levels of assignment (primary, secondary and/or tertiary) to be used in the page. For a call-related page, the default page message shall include information about the call, including the call type and room number requiring assistance. The attendant may edit the message or make use of a pre-programmed list of patient needs.

Automatic paging shall be governed by paging assignments. When a call is placed from a patient station, the system shall automatically page pagers or pager groups assigned to that station or its zones. The properties of each assignment shall determine its applicability to the page, ensuring that pages reach exactly the staff members they concern. For example, a call of a particular type shall activate only those assignments whose call filters allow that type. For each call, automatic paging shall begin at the primary level and repeat at regular intervals. If the call is not reset, the page shall escalate to include up to 8 levels. The pager shall display the location where the call originated (room ID), as well as the call type. Call resets and faults may also be paged automatically.

Systems that require additional software or licensing to enable paging shall not be acceptable.

The transmitter shall be TekTone® NC365B

The alphanumeric pager shall be TekTone® NC397A

The LED messaging sign shall be TekTone® SI005

## Email Output

The contractor shall provide:

An interface to the facility’s mail server, model LS453 Email Output software that routes calls from the nurse call system to email addresses as required.

The email output software shall enable real-time event notifications to be sent to email addresses and any device with an email gateway.

The LS453 Email Output software shall allow users to create scheduled reports that are emailed directly to the specified users.

Email output software shall be available as an optional feature. It is not included with the basic nurse call system.

## Wireless Nurse Call Interface

The contractor shall provide, as required;

An interface to the Tek-CARE®500 wireless nurse call system, model LS500. The LS500 interface shall enable the nurse call system server to:

- Display and respond to calls from Tek-CARE®500 wireless transmitters.

- Allow a single event monitor to display system events for both systems.

- Allow a single reporting system to handle reporting for both systems.

- Use one pocket paging interface to handle both systems.

The wireless nurse call interface shall be available as an optional feature. It is not included with the basic nurse call system.

## Apple iOS Integration

The contractor shall provide, as required, an interface to up to ten (10) Apple iPod Touch® and iPhone® and/or up to ten (10) Apple TV® devices. The interface for iPod Touch® and iPhone® shall be TekTone® LS621. The Apple TV® interface shall be TekTone® LS622-series. The required app server shall be TekTone® LS620.

The iOS® integration shall display calls from the Nurse Call System to improve staff efficiency and awareness. Mobile iOS® clients shall include the ability to set up to seven (7) individual statuses for each station present on the system. The iOS® integrations shall only function when connected to the facility’s wireless LAN. iOS® integrations that allow access to system calls when the iOS® device is not on facility premises shall not be accepted.

Apple TV® devices and Mobile iOS® client assignments shall be configurable in the same way as the radio pagers can be.  Messaging shall be automatic and support manual messaging.

## Foreign Systems Integration

The contractor shall:

Provide, as required, a Tek-BRIDGE® Server with LS600 Tek-ALERT® software and appropriate connection hardware to enable the Nurse Call System to monitor and display calls and status of connected foreign systems such as security systems, fire alarm systems, and more.

The monitored foreign system must supply an accessible data stream via one of the following protocols: Serial (event printing, Scope paging, ContactID), COMPII, TAP, HL7, TCP/IP, TCP/IP Listener or UDP.  Contact TekTone® if a protocol is not listed, as protocol libraries will grow.

Foreign Systems Integration shall support the following features

- Once properly configured, the Tek-ALERT® Integration Manager shall receive events from monitored foreign systems and annunciate the events on the Nurse Call System.

- Events from monitored foreign systems shall appear in the reports created by the Tek-CARE® Reporting System. These events may also be sent to pagers/apps as configured in assignment settings.

- The Tek-ALERT Integration Manager shall not be used as the primary annunciation point for monitored foreign systems, but shall function as a secondary annunciation point. Tek-ALERT programming and use shall not interfere with the primary annunciation settings of the monitored foreign system(s).

## Media Gateway Server

The contractor shall provide, as required;

A Media Gateway, model NC465, to provide call and event notifications to facility VoIP extensions and deliver pertinent information on the telephone display.  It shall provide protocol conversion between TekTone’s audio manager and a facility PBX that uses SIP G.711 as the primary protocol.

The NC465 also shall provide two-way voice communication between patient room stations and facility VoIP extensions.

The Media Gateway, with its 2 network interfaces, shall be capable of connecting to the both the PBX network and Tek-CARE® network and shall be operationally transparent to the nurse call system.

The interface shall support the following features:

- Provides up to 36 audio paths

- Display of patient or resident nurse call alarms

- Accept incoming voice calls from the nurse call system

- Reset a nurse call alarm (low priority only and can be enabled or disabled)

- In-call options (connected call)

  - Place a Level 1, 2 or 3 service request

  - Place a STAT call

- Place a call to a room station, zone or master

- Receive manual text messages from a master or remote monitor

- Software updates along with any other nurse call system module update

Telephone extension assignments shall be configurable in the same way as the radio pagers can be and support manual messaging.

Backup power for the integration server shall be supplied by an uninterruptable power supply (UPS). The UPS shall be TekTone® PK250B.

## ADT Interface Using the HL7 Standard Software

The contractor shall:

    - Provide an HL7 (v. 2.3) compliant interface for the purpose of receiving relevant patient information from the facility’s Admit, Discharge, Transfer system. Relevant patient information shall include patient name, room number, bed number and doctor name.

The interface shall:

- Insert standard ADT field information into nurse call system fields.

- Update in real time.

## System Diagnostics

A.	All Network modules, Nurse Master Stations and IR400-series stations in the system shall be continuously supervised for power and data. Each IR400-series station shall in turn monitor its peripheral bus devices and dome lamps as previously specified. All system faults shall be annunciated on a Nurse Master Station with pertinent information as to the type and location of fault.

B.	Provide separate Chassis Short Monitor, Model NC455, to monitor integrity of field wiring for short(s) to earth ground. An audible alert tone shall sound if short is detected. Provide test button to test that all functions are operational.

C.	Provide separate Power Fault Monitor, Model NC460, to notify staff of loss of power. Should a loss of power or power brownout occur, an audible alarm shall sound. Provide test button to test functional operation.

## Power Monitor and Supervised Code System Alarms

The contractor shall:

- Provide separate Power Fault Monitor, Model NC460, to notify staff of loss of power. Should a loss of power or power brownout occur, an audible alarm shall sound. Provide test button to test functional operation.

- Provide, as required, Chassis Short Monitor Model NC455 and Code Call Switches to qualify system for Supervised Code Blue operation. As such, the system shall supervise code blue stations for power, data, and faults caused by open or shorted filaments, or by missing code blue bulbs.

- Provide power supervision of PM453 Master/Station Modules. If power is switched to battery backup, master stations shall display “Battery Backup Active” (default text) once voltage has dropped to approximately 24VDC.

- Systems that do not provide all of the functions of this Section are not acceptable.

## Part 3: Execution

## Supervision

Only TekTone® factory-certified installers shall install, service, and maintain the specified network system.

The manufacturer shall have the equipment manufacturer’s engineer or their designated agent inspect the installation and operation of this network to determine that the network complies with all standards listed in Section 2.3.

## In-Service Training

The contractor shall provide thorough training of all nursing staff assigned to those nursing units receiving newly networked nurse/patient communications equipment. This training shall be developed and implemented to address two different types of staff: floor nurses/staff shall receive training that is specific to their tasks and responsibilities, and similarly, unit secretaries (or any person whose specific responsibilities include answering patient calls and dispatching staff) shall receive operational training from their perspective. A separate training room shall be set up that allows this type of individualized training utilizing in-service training unit, prior to the turning over of the new system.

## Wiring

A.	The contractor shall terminate all wiring with manufacturer’s approved connectors. The use of wire nuts is prohibited.

B.	All wiring shall be free from shorts and faults. Wiring shall be UL® Listed, and installed in accordance with ANSI/NFPA 70, Article 25 and applicable sections of ANSI/NFPA 99, and compliant with the manufacturer’s installation and maintenance manual and specifications.

C.	Nurse patient communications network wiring shall not be run in the same conduit with other systems (e.g.: Class 1 AC power distribution, fire alarm, entertainment systems, lighting controls, etc.).

## Electrical Power Connections

A.	It shall be the responsibility of the facility to provide a dedicated 120 VAC, 60 Hz conduit feed into the equipment cabinet. This power feed shall not have any other devices connected to it. A 20-amp circuit breaker shall be located in the electrical sub-panel. This circuit breaker shall be labeled “Nurse Call” along with identification of the Nursing Unit that shall be controlled by this circuit breaker. This electrical circuit shall be connected to the facility’s emergency power system for automatic power switchover during loss of utility power.

B.	Connect all network system power supplies and equipment cabinets to a common earth ground approved for the application utilizing a 14 AWG or larger solid conductor which is at minimum the same conductor size as the AC feed wires.

## Environmental Protection

The installing contractor shall make certain that all central equipment is accessible for service. Contractor shall notify specifying authority if designated equipment closet does not meet manufacturer’s requirements for heat, radiation or static electricity.

## Protection of Network Devices

The contractor shall protect network devices during unpacking and installation by wearing manufacturer-approved ESD wrist straps tied to earth ground. The wrist strap shall meet OSHA requirements for prevention of electrical shock if technician comes in contact with high voltage.

## Cleaning and Patching

It shall be the responsibility of the contractor to keep the work area clear of debris and to clean area daily at completion of work.

It shall be the responsibility of the contractor to patch and paint any wall or surface that has been disturbed by the execution of this work.

## Drawings

Provide as-built drawings of all installed network components and associated wiring on building plans. Final payment for work will not be authorized unless these drawings are supplied.

## Website Disclaimer and Legal Notice for Downloadable 
Architects & Engineers Specifications

The specifications attached herein are being supplied in a text format for your convenience. TekTone® Sound & Signal Mfg., Inc. is not responsible for any changes made to the original documents.

As with any other TekTone® Sound & Signal Mfg., Inc. document, please be advised that the information was prepared by TekTone® Sound & Signal Mfg., Inc. as general information and a service to our customers. While we endeavor to supply accurate information, we can take no responsibility for errors or misinterpretations in any of the content of this site. We also reserve the right to remove or change literature and pages from this site where legal requirements are contravened or where complaints are received and upheld.

Except in the case of third-party material reproduced on this site, the information and all content, including logos and artwork, provided herein is protected under the U.S. and Foreign Copyright Laws and may not be copied, reproduced or retransmitted by any means, electronically or otherwise, without the express written permission of TekTone® Sound & Signal Mfg., Inc.

Modifications. All modified specifications will include the following ledger: Modified specifications are based upon TekTone® Sound & Signal Mfg., Inc. product specifications. Modified specifications have not been reviewed or approved by TekTone® Sound & Signal Mfg., Inc., unless so noted.

Indemnification. It is agreed that TekTone® will be held harmless from any claims, causes of action, pending or threatened, losses, awards, damages or attorney’s fees arising out of any suit or claim relating to the use and/or modification of the specifications.

Reservation of Rights. This disclosure is not intended, nor shall it be construed, to release any right, title or interest in any name, trademark, trade name, advertising slogan, specification, or confidential information, but is solely intended to provide a revocable license to use the specification material for the contemplated use.

END OF SECTION